HOW TO USE RS-232 SIMULATOR PROGRAM

- Please first check ECR Comm. Type of PD-II.

- Please select TYPE-4 for ECR comm. Type.

ECR Comm. Type 을 TYPE-4 로 설정 합니다.

 Please kindly refer the service manual.

- Please check RS-232C cable of the scale and PC whether they are connected properly.

 Pin Connection (RS-232 Standard : 2,3,5 Pin)

 For the detail information, please refer the service manual.

- Please execute “RS232C SIMULATOR” program.

- Please check the PORT from the SETTING display.

- Please check “Parity, Data, Stop Bit”.

- Please check BPS.

The setting No. 6.7 of PD-II for Type-4 communication  Settings value is as below;

Parity bit : Even Parity.   Setting value = E

Data bit  :  7 data bit     Setting value = 7

Stop bit  :  1 Stop bit     Setting value = 1

BPS      :  9600 BPS     Setting value = 9600

- Please select “Message Data” and then, HEXA. The rest of the setting will remain the same as below;

- In the DATA SEND, please put the value “57” which is the communication command value for the Type 4 and then, press ENTER KEY.

- In the DATA SEND, please input again  “0D” and then, press “ENTER” KEY.

 “W” ASCII value = HEXA value is “57”

 “CR” ASCII = HEXA Value = “0D”

- If the display appear as below, it means that the scale and PC is communicating properly.